Załącznik A do Formularza oferty – język angielski

FUNCTIONS, TECHNICAL PARAMETERS AND REQUIRED CONDITIONS

Full name of the device, quantity: System for testing mitochondrial functions – 1 pc

Type, model: .......................................................................................

Manufacturer: .......................................................................................

| No. | Parameters and required conditions: | Description of the parameters and conditions offered (to be completed by the Contractor): |
| --- | --- | --- |
|  | Two-chamber respirometer for monitoring oxygen consumption or production using small amounts of biological sample (mg) |  |
|  | Stainless steel housing |  |
|  | Equipped with the electronic Peltier thermostat |  |
|  | The temperature in both chambers of the respirometer can be adjusted with stability of min. ± 0.01 °C in the range of at least 5 to 45 °C |  |
|  | Containing two independently controlled electromagnetic stirrer systems |  |
|  | Adjustable mixing speed in the range from min. 200 to 800 rpm |  |
|  | Ability to monitor barometric pressure with minimal resolution of ±0.1 kPa |  |
|  | Design and built using specialized materials, avoiding the use of oxygen-absorbing plastics (eg. Teflon) |  |
|  | Limit of detection of O2 flux is below 1.5 pmol O2 s-1 mL-1 |  |
|  | Resolution of oxygen concentration: 1000 µM to 0.005 µM. The signal noise (standard deviation) at zero oxygen concentration is <0.05 μM O2 |  |
|  | High signal linearity range, allowing display of minimum respiratory rates in the full range of oxygen (from air saturation to zero oxygen) |  |
|  | Equipped with the software-controlled titration-injection micropump that allows to perform pre-programmed automatic titrations, and steady-state injections. Equipped with feedback-controlled maintenance of oxygen levels, pH, or other parameters in the respirometer chambers |  |
|  | Equipped with dedicated software for high-resolution respirometry testing |  |
|  | Equipped with the module for simultaneous measurement of oxygen consumption and NAD(P)H autofluorescence |  |
|  | Equipped with the module for measuring the Q-redox state and cyclic voltammetry |  |
|  | Equipped with the fluorometric module for measurement of hydrogen peroxide production, ATP production, mitochondrial membrane potential, Ca2+ |  |
|  | Equipped with the module of small volume chambers to reduce the sample amount needed for oxidative phosphorylation analysis and mitochondrial respiratory function (chambers with the operation volume – at least of 0.6 mL) |  |
|  | Equipped with integrated computer with touchscreen |  |
|  | Main unit weight up to 15 kg |  |
|  | Maximum dimensions of the main unit 50x35x30 cm (LxWxH) |  |

The above functions and parameters are the minimum required conditions. Failure to meet them shall result in rejection of the offer.

The Economic Operator shall fill in column 3 of the table with the technical and functional parameters, i.e. functions, technical parameters and conditions of the models offered, as appropriate. The Economic Operator is required to specify the technical and functional parameter in the units indicated in column 2. The Contracting Entity allows the word YES or equivalent to be provided if the Contracting Entity is able to verify the parameter in the qualitative means of proof submitted with the tender offer. The Contracting Entity shall allow the subject of the public contract with parameters better than required.

Should the Economic Operator fail to fill in any row of column 3, the Contracting Entity shall verify this parameter in the subjective means of proof and assume that the parameter described in the subjective means of proof has been offered.